UV-GUARD

Autonomous UV Disinfection & Environment Monitoring System

for OT / ICU Medical Environments

COMPLETE PROJECT ENGINEERING DOCUMENT

Hardware · Software · Firmware · Integration · Design · Deployment

| Team 1 | Team 2 | Integration | Timeline |
| --- | --- | --- | --- |
| Environment Monitoring | Disinfection Robot | System Integration | 3–4 Months To Market |

Version 1.0  |  April 2026  |  CONFIDENTIAL – PROJECT TRACKING DOCUMENT

# TABLE OF CONTENTS

1.  Project Overview & Architecture	3

2.  System Safety Logic & Interlock Protocol	5

3.  Team 1 – Environment Monitoring System	7

4.  Team 2 – Disinfection Mobile Robot	12

5.  Integration Architecture	20

6.  Software & Firmware Development Guide	23

7.  Hardware BOM & Procurement	28

8.  PCB & Enclosure Design	31

9.  Medical Standards & Compliance	33

10.  Testing & Validation Protocol	35

11.  Level-1 Project Tracking Report (3–4 Month Roadmap)	38

12.  Budget Estimation	43

# 1. PROJECT OVERVIEW & SYSTEM ARCHITECTURE

## 1.1 Project Name & Objective

UV-GUARD is an autonomous, safety-interlocked UV-C disinfection robot integrated with a real-time environment monitoring station, purpose-built for Operating Theatres (OT) and Intensive Care Units (ICU). The system ensures zero human exposure during UV disinfection, complete spatial coverage through AI-guided navigation, and continuous post-disinfection environment quality reporting aligned with WHO, NABH, and ASHRAE standards.

## 1.2 System Sub-Systems

| Sub-System | Function | Responsible Team | Status |
| --- | --- | --- | --- |
| SS-01 | Thermal Human Detection | Team 2 (Robot) | Design Phase |
| SS-02 | Door & Window Sensor Interlock | Team 2 (Robot) | Design Phase |
| SS-03 | UV-C Disinfection Module | Team 2 (Robot) | Design Phase |
| SS-04 | Autonomous Navigation & Coverage | Team 2 (Robot) | Design Phase |
| SS-05 | Real-time Organism Detection Camera | Team 2 (Robot) | Design Phase |
| SS-06 | Environment Sensor Array | Team 1 (Monitoring) | Design Phase |
| SS-07 | Display Dashboard (HMI) | Team 1 (Monitoring) | Design Phase |
| SS-08 | Central Control & Safety Logic | Integration Team | Design Phase |
| SS-09 | Reporting & Data Logging | Integration Team | Design Phase |

## 1.3 High-Level System Architecture

Master Controller (Raspberry Pi 5 / Jetson Orin Nano): Central brain orchestrating all sub-systems. Runs safety state machine, navigation AI, computer vision pipeline, and MQTT communication bus.

- Safety Interlock Layer → Thermal cameras + door/window sensors feed safety FSM

- Disinfection Layer → UV-C lamp control + robot drive system

- Perception Layer → Camera feeds YOLOv8 for organism detection + SLAM for navigation

- Environment Layer → PMS5003, BME280, CO2, VOC sensors → HMI dashboard

- Reporting Layer → SD card + WiFi → cloud/mobile app log

# 2. SYSTEM SAFETY LOGIC & INTERLOCK PROTOCOL

## 2.1 Safety State Machine (FSM)

UV-GUARD operates a strict 7-state Finite State Machine. UV-C lamps CANNOT energize unless ALL safety conditions are simultaneously TRUE. A single breach triggers immediate lamp shutdown.

| State | Name | Condition to Enter | Action |
| --- | --- | --- | --- |
| S0 | IDLE | System power-on default | All outputs OFF. Await start command. |
| S1 | PRE-CHECK | Start command received | Run thermal scan + check all door/window sensors. |
| S2 | ARMED | No humans detected + all doors/windows CLOSED | Alert via buzzer/light. 30-sec countdown before UV ON. |
| S3 | DISINFECTING | All S2 conditions still TRUE after countdown | UV-C ON. Robot begins coverage pattern. Continuous monitoring. |
| S4 | BREACH DETECTED | Any sensor fails during S3 | UV-C OFF immediately (<50ms). Sound alarm. Log event. |
| S5 | RE-CHECK | Breach cleared + operator confirmation | Return to S1 pre-check cycle. |
| S6 | COMPLETE | Coverage map 100% + timer elapsed | UV-C OFF. Post-scan mode begins. Generate report. |

## 2.2 Sensor Interlock Conditions

### Thermal Human Detection

- Sensor: MLX90640 32×24 Thermal Camera Array (×2 units for full room coverage)

- Detection threshold: Any pixel >32°C within defined ROI = HUMAN DETECTED

- Scan frequency: Every 500ms during ARMED and DISINFECTING states

- False-positive handling: 3 consecutive positive readings required before breach trigger (1.5 sec)

### Door & Window Sensors

- Type: Magnetic Reed Switches (MC-38) on all OT/ICU doors + windows

- Logic: NC (Normally Closed) = SAFE. Any OPEN = BREACH

- Response time: <10ms interrupt-driven on microcontroller GPIO

- Wiring: Supervised loop with EOL (End of Line) resistor to detect tampering

### UV-C Interlock Hardware

- Hardware relay with normally-open contacts – relay de-energizes = UV OFF (fail-safe)

- Separate hardware watchdog timer: if MCU freezes → relay trips within 2 seconds

- Emergency Stop button (red mushroom head) on robot + wall-mounted panel

## 2.3 Safety Wiring Diagram (Description)

Power → Main MCB → Safety Relay Chain: [Thermal OK] AND [All Doors Closed] AND [Watchdog Pulse] AND [No E-STOP] → UV Lamp Driver Board. Any link breaks = UV OFF.

# 3. TEAM 1 – ENVIRONMENT MONITORING SYSTEM

## 3.1 Sensor Selection & Justification

| Sensor | Parameter | Range | Accuracy | Interface |
| --- | --- | --- | --- | --- |
| PMS5003 | PM1.0, PM2.5, PM10 (Particulate Matter) | 0–500 µg/m³ | ±10 µg/m³ | UART (9600 baud) |
| BME280 | Temperature, Humidity, Barometric Pressure | T: -40–85°C H: 0–100%RH P: 300–1100 hPa | T: ±0.5°C H: ±3% | I2C / SPI |
| MH-Z19B | CO2 Concentration | 0–5000 ppm | ±50 ppm ±5% | UART / PWM |
| SGP30 | TVOC, eCO2 | TVOC: 0–60000 ppb | ±15% of reading | I2C |
| MLX90614 | Spot IR Temperature (surface sterilization check) | −70–382°C | ±0.5°C | I2C (SMBus) |
| BH1750 | Ambient Light (lux) (UV verification) | 1–65535 lux | ±20% | I2C |

## 3.2 Medical OT/ICU Environmental Standards

The HMI display shall classify each parameter using the following medically-referenced thresholds (WHO, ASHRAE 170, NABH, and CDC OT standards):

| Parameter | Unit | EXCELLENT ✅ | GOOD ✅ | WARNING ⚠️ | DANGER 🔴 |
| --- | --- | --- | --- | --- | --- |
| Temperature | °C | 20–21 | 18–24 | 16–17 or 25–26 | <16 or >26 |
| Relative Humidity | %RH | 50–55 | 45–60 | 35–44 or 61–70 | <35 or >70 |
| CO2 | ppm | <600 | 600–800 | 800–1000 | >1000 |
| PM2.5 | µg/m³ | <5 | 5–10 | 10–25 | >25 |
| PM10 | µg/m³ | <10 | 10–20 | 20–50 | >50 |
| TVOC | ppb | <50 | 50–200 | 200–500 | >500 |
| Air Changes/Hr | ACH | >25 | 20–25 | 15–20 | <15 |
| OT Pressure | Pa (vs corridor) | +8 to +12 | +5 to +15 | +3 to +5 | <+3 |

## 3.3 HMI Display Design

### Hardware

- Primary Display: 10.1" HDMI IPS Touchscreen (1280×800) mounted on wall bracket

- Secondary Display: 2.8" SPI TFT (ILI9341) on robot body for local status

- Buzzer: 5V active piezo for audio alerts

- RGB LED Strip: WS2812B addressable LEDs (Green/Yellow/Red ambient status)

### Software Stack (Python + Tkinter/PyQt5)

- Main dashboard: Real-time gauges for all 6+ sensor parameters

- Color-coded status badges: EXCELLENT (green), GOOD (teal), WARNING (amber), DANGER (red)

- Trend charts: 1-hour rolling chart using matplotlib embedded

- Infection Risk Score: Composite score 0–100 calculated from weighted sensor values

- Surgery Readiness Banner: 'READY FOR SURGERY ✓', 'CONDITIONAL ⚠', 'NOT SAFE ✗'

## 3.4 Team 1 – Firmware Architecture

### Microcontroller: Raspberry Pi 4 Model B (4GB) + Arduino Mega 2560 (sensor co-processor)

Arduino Mega handles time-critical sensor polling and sends data via USB-Serial to RPi4 at 115200 baud. RPi4 runs Python application for data processing, display, logging, and MQTT publishing.

| Layer | Technology | Responsibility |
| --- | --- | --- |
| Sensor Layer | Arduino Mega (C++) | Poll PMS5003, BME280, MH-Z19B, SGP30 every 2 sec. Send JSON over Serial. |
| Data Processing | Python 3.11 (RPi4) | Parse JSON, calculate composite scores, classify per medical thresholds. |
| Display Layer | PyQt5 / Tkinter | Render dashboard. Update every 5 sec. Touch input for navigation. |
| Storage | InfluxDB (local) | Time-series storage. Retain 90-day history. |
| Communication | MQTT (Mosquitto) | Publish sensor data to broker for robot + cloud integration. |
| Reporting | Python + ReportLab | Auto-generate PDF disinfection completion reports. |

## 3.5 Team 1 Code Modules

### Module 1: sensor_manager.py

Purpose: Unified sensor abstraction layer. Each sensor has its own class inheriting from BaseSensor.

class PMS5003Sensor(BaseSensor):
    def read(self) -> dict:
        # Parse 32-byte frame from UART
        # Returns: {pm1: float, pm2_5: float, pm10: float}
class BME280Sensor(BaseSensor):
    def read(self) -> dict:
        # Read via I2C using smbus2
        # Returns: {temp: float, humidity: float, pressure: float}
class CO2Sensor(BaseSensor):
    def read(self) -> dict:
        # Send 0xFF 0x01 0x86 command, parse 9-byte response
        # Returns: {co2_ppm: int}

### Module 2: classifier.py

Purpose: Map raw sensor readings to medical classification levels and compute Surgery Readiness Score.

THRESHOLDS = {
    'temp':     [(20,21,'EXCELLENT'), (18,24,'GOOD'), (16,26,'WARNING')],
    'humidity': [(50,55,'EXCELLENT'), (45,60,'GOOD'), (35,70,'WARNING')],
    'co2':      [(0,600,'EXCELLENT'), (600,800,'GOOD'), (800,1000,'WARNING')],
    'pm2_5':    [(0,5,'EXCELLENT'),  (5,10,'GOOD'),   (10,25,'WARNING')],
}
def classify(param, value) -> str:
    for low, high, label in THRESHOLDS[param]:
        if low <= value <= high: return label
    return 'DANGER'
def surgery_readiness_score(readings) -> int:
    # Weighted composite: temp(20%) + humidity(20%) +
    # co2(25%) + pm2_5(25%) + tvoc(10%) -> 0-100

# 4. TEAM 2 – DISINFECTION MOBILE ROBOT

## 4.1 Robot Mechanical Design

### Chassis Specification

| Parameter | Specification |
| --- | --- |
| Chassis Type | Differential Drive 4-Wheel (2 driven + 2 casters) |
| Overall Dimensions | 600mm (L) × 500mm (W) × 1400mm (H) including UV tower |
| Material | Aluminum extrusion frame (20×20 V-slot) + ABS enclosures |
| Drive Motors | 2× 12V 30RPM Planetary Gear DC Motor with encoder (200 PPR) |
| Wheels | 150mm diameter polyurethane rubber wheels |
| Payload Capacity | 15 kg (for electronics + UV lamps + battery) |
| Speed | 0.2 m/s operating speed (slow for thorough disinfection) |
| Battery | 24V 20Ah LiFePO4 (lithium iron phosphate – safe for medical) |
| Runtime | 4–6 hours per charge (covers 50–60 m² OT) |
| Charging | Auto-docking charging station at OT entrance |

## 4.2 UV-C Lamp System

### Lamp Selection

- Type: Amalgam UV-C Germicidal Lamps (254 nm wavelength – peak germicidal)

- Power: 2× 55W lamps on rotating arm + 1× 36W lower lamp for floor/corner

- UV Intensity: Minimum 25 mW/cm² at 1 meter distance (WHO standard for 3-log reduction)

- Lamp Life: 9,000 hours (replace annually)

- Warm-up Time: 3–5 minutes to full intensity (system accounts for warm-up in FSM)

- Driver: Electronic ballast with 0-10V dimming for controlled ramp-up

### Coverage Algorithm

- Room map imported from CAD floor plan OR auto-generated by SLAM on first run

- Coverage Path: Modified Boustrophedon (lawnmower) pattern with 0.5m pass spacing

- Corner Strategy: Robot pauses 45 seconds at each corner, rotates 360° for full exposure

- Dose Calculation: UV dose = Intensity (mW/cm²) × Time (sec). Target: ≥4000 mJ/cm² for C. difficile

- UV dose sensor (radiometer) on robot confirms actual delivered dose at each position

## 4.3 Navigation System

### SLAM (Simultaneous Localization and Mapping)

- LiDAR: RPLIDAR A1M8 (360°, 12m range, 8000 samples/sec) for SLAM mapping

- SLAM Algorithm: GMapping or Cartographer ROS2 package

- Localization: AMCL (Adaptive Monte Carlo Localization) during operation

- Path Planner: Nav2 (Navigation2) ROS2 stack with DWA local planner

- Odometry: Motor encoders + MPU6050 IMU fusion (Kalman filter)

- Obstacle Avoidance: 3× HC-SR04 ultrasonic sensors (front, left-front, right-front)

### Object Detection for Coverage Verification

- Camera: Intel RealSense D435i (RGB-D) – provides depth + color

- Model: YOLOv8n-seg (nano, real-time) trained on medical equipment classes

- Purpose: Detect furniture/equipment to navigate around AND verify UV shadow zones

- Shadow Zone Handling: If equipment blocks UV, robot moves to alternate angle and re-exposes

## 4.4 Thermal Scanning System (Anti-Human Safety)

### Hardware

- Sensor: 2× MLX90640 Thermal Camera (32×24 pixels, 110° FoV, I2C)

- Mounted: Front-top and rear-top of robot for 270° coverage

- Fixed wall unit: 1× MLX90640 at door for pre-entry scan (ceiling mount)

### Detection Algorithm (Python/C++)

def detect_human(thermal_frame) -> bool:
    # 1. Apply 5x5 median filter (noise reduction)
    # 2. Threshold: pixels > 32.0°C flagged as ROI
    # 3. Connected components analysis (min 50 pixels = person-sized)
    # 4. Shape analysis: aspect ratio 1.5-4.0 (human body)
    # 5. Return True if blob found with 3 consecutive frames
    # False positive guard: exclude equipment with fixed thermal signature
    hot_pixels = frame[frame > 32.0]
    if len(hot_pixels) > 50:
        blobs = connected_components(frame > 32.0)
        for blob in blobs:
            if 1.5 < blob.aspect_ratio < 4.0:
                return True
    return False

## 4.5 UV Fluorescence Imaging System (Organism Detection)

### Concept

Under UV-A (365nm) illumination in a darkened room (blackout condition after UV-C cycle), biological materials fluoresce. Blood stains fluoresce red-pink, bacteria colonies show as blue-white, fungi show green-yellow, and tissue particles show orange. This allows post-disinfection verification.

### Hardware for UV Fluorescence Scanning

- UV-A LED Panel: 365nm UV-A (blacklight) – 20W panel mounted on robot arm

- Camera: Raspberry Pi HQ Camera Module with 16mm lens (no IR filter – sensitive to fluorescence)

- Optical Filter: 400nm long-pass optical filter on camera lens (blocks UV, passes visible fluorescence)

- Room condition: OT blackout blinds + door closed (pre-programmed blackout routine)

### Detection Pipeline

- Capture high-res image under UV-A in darkened room

- Color segmentation: HSV thresholding for red (blood), green-yellow (fungi), blue-white (bacteria)

- CNN classifier (MobileNetV2 fine-tuned) for 4 classes: Clean, Blood, Bacteria/Fungi, Tissue

- Overlay detection map on floor plan and display on HMI with GPS-tagged location

- Risk score: % contaminated area vs total floor area

## 4.6 Robot Electronics Architecture

| Component | Module | Function |
| --- | --- | --- |
| Main Computer | NVIDIA Jetson Orin Nano 8GB | ROS2, SLAM, YOLOv8, Computer Vision pipeline |
| Safety MCU | STM32F4 Discovery | FSM, relay control, watchdog, real-time sensor polling |
| Motor Controller | L298N × 2 or ODrive | Dual motor PWM control with encoder feedback |
| Thermal Cameras | MLX90640 × 2 (I2C) | Human detection – safety interlock |
| LiDAR | RPLIDAR A1M8 (USB) | SLAM mapping and navigation |
| Depth Camera | Intel RealSense D435i (USB3) | 3D perception, object detection |
| UV-C Control | Safety relay + MOSFET driver | Controlled energization of UV-C lamps |
| UV-A LED Driver | Constant current LED driver | Fluorescence imaging illumination |
| IMU | MPU6050 (I2C) | Orientation, wheel slip detection |
| Door Sensors | MC-38 Reed switches × 6 | All doors + windows supervised loop |
| Communication | WiFi (RPi) + CAN bus | Inter-board + cloud communication |
| Power Management | 24V→5V, 12V DC-DC converters | Regulated power for all boards |

# 5. INTEGRATION ARCHITECTURE

## 5.1 Communication Protocol

All subsystems communicate via MQTT protocol over local WiFi network. A Mosquitto broker runs on Team 1's Raspberry Pi 4 which also serves as the local network gateway.

| MQTT Topic | Publisher | Subscriber / Purpose |
| --- | --- | --- |
| uvguard/environment/sensors | Team 1 RPi4 | Robot + Cloud: environmental readings |
| uvguard/safety/human_detected | Robot STM32 | All systems: emergency halt trigger |
| uvguard/safety/door_status | Robot STM32 | All systems: door/window state |
| uvguard/robot/state | Robot Jetson | HMI: current FSM state display |
| uvguard/robot/coverage_map | Robot Jetson | HMI: real-time coverage progress |
| uvguard/uvc/control | Safety MCU | Lamp driver: ON/OFF/DIM commands |
| uvguard/detection/organisms | Robot Jetson | HMI + Report: fluorescence scan results |
| uvguard/report/complete | Jetson | Cloud + Print: final disinfection report |

## 5.2 Data Flow Diagram (Described)

Sensors → Arduino/STM32 reads raw sensor data (2sec cycle)

Processing → Raspberry Pi classifies + scores data

Broker → MQTT Mosquitto distributes data to HMI, Robot, Cloud

Display → PyQt5 dashboard updates every 5 seconds on 10" screen

## 5.3 API Endpoints (REST for Cloud/Mobile)

- GET  /api/v1/environment/latest  → Latest sensor reading JSON

- GET  /api/v1/disinfection/history → Past 30 sessions log

- POST /api/v1/disinfection/start   → Remote start disinfection (requires auth token)

- GET  /api/v1/report/{session_id}  → Download PDF report

- WS   /ws/live                     → WebSocket for real-time dashboard on mobile

# 6. SOFTWARE & FIRMWARE DEVELOPMENT GUIDE

## 6.1 Development Environment Setup

### Team 1 (Environment Monitoring) – Raspberry Pi 4

- OS: Raspberry Pi OS Lite 64-bit (Bookworm)

- Python 3.11 + virtual environment: python3 -m venv uvguard_env

- Libraries: pip install pyserial smbus2 paho-mqtt influxdb-client PyQt5 matplotlib reportlab

- InfluxDB 2.0: Install via apt. Create bucket 'uvguard_sensors'.

- Mosquitto MQTT: sudo apt install mosquitto mosquitto-clients

### Team 2 (Robot) – Jetson Orin Nano

- OS: JetPack 6.0 (Ubuntu 22.04 LTS based)

- ROS2 Humble: Follow official ROS2 Humble install guide for ARM64

- Install: ros-humble-navigation2, ros-humble-slam-toolbox, ros-humble-rplidar-ros

- Python: pip install ultralytics opencv-python numpy scipy paho-mqtt pyrealsense2

- YOLOv8: pip install ultralytics  (use yolov8n.pt base, fine-tune on medical equipment dataset)

## 6.2 ROS2 Package Structure

uvguard_ws/
├── src/
│   ├── uvguard_robot/          # Main robot package
│   │   ├── uvguard_robot/
│   │   │   ├── safety_fsm.py       # Finite State Machine
│   │   │   ├── thermal_detector.py # Human detection
│   │   │   ├── uvc_controller.py   # UV lamp control
│   │   │   ├── coverage_planner.py # Path planning
│   │   │   └── organism_detector.py# UV-A scan pipeline
│   │   ├── config/
│   │   │   ├── nav2_params.yaml
│   │   │   └── slam_params.yaml
│   │   └── launch/
│   │       ├── robot_launch.py
│   │       └── disinfection_launch.py
│   └── uvguard_msgs/           # Custom ROS2 messages
│       └── msg/
│           ├── SafetyStatus.msg
│           ├── CoverageMap.msg
│           └── OrganismDetection.msg

## 6.3 Safety FSM Implementation (Key Code)

class SafetyFSM(Node):
    STATES = ['IDLE','PRE_CHECK','ARMED','DISINFECTING','BREACH','RECHECK','COMPLETE']
    def transition(self, event):
        if self.state == 'ARMED' and event == 'UV_START':
            if self.all_clear():  # All safety conditions TRUE
                self.state = 'DISINFECTING'
                self.activate_uvc()  # Energize relay
        elif self.state == 'DISINFECTING':
            if event in ['HUMAN_DETECTED', 'DOOR_OPEN', 'WATCHDOG_FAIL']:
                self.deactivate_uvc()  # IMMEDIATE shutdown
                self.state = 'BREACH'
                self.publish_alarm(event)
    def all_clear(self) -> bool:
        return (not self.human_present and
                self.all_doors_closed and
                self.all_windows_closed and
                not self.estop_pressed)

## 6.4 YOLOv8 Training for Organism Detection

- Base model: YOLOv8n-seg (instance segmentation)

- Dataset: Collect 500+ UV-A fluorescence images per class (Clean, Blood, Bacteria, Fungi, Tissue)

- Annotation: Use Roboflow for bounding box + polygon annotation

- Training: yolo train model=yolov8n-seg.pt data=organisms.yaml epochs=100 imgsz=640

- Validation: 80/10/10 train/val/test split. Target mAP50 >0.75 before deployment

- Deployment: Export to TensorRT FP16 for Jetson acceleration: yolo export format=engine half=True

# 7. HARDWARE BILL OF MATERIALS (BOM)

## 7.1 Team 1 – Environment Monitoring BOM

| # | Component | Model | Qty | Unit Price | Total | Source |
| --- | --- | --- | --- | --- | --- | --- |
| 1 | Single Board Computer | Raspberry Pi 4B 4GB | 1 | ₹6,000 | ₹6,000 | Robu.in |
| 2 | Arduino Mega | Arduino Mega 2560 | 1 | ₹1,500 | ₹1,500 | Robu.in |
| 3 | Air Quality Sensor | PMS5003 | 1 | ₹1,800 | ₹1,800 | Amazon |
| 4 | Temp/Humidity/Pressure | BME280 Module | 2 | ₹350 | ₹700 | Robu.in |
| 5 | CO2 Sensor | MH-Z19B | 1 | ₹2,200 | ₹2,200 | Amazon |
| 6 | VOC Sensor | SGP30 | 1 | ₹800 | ₹800 | Robu.in |
| 7 | IR Temperature | MLX90614 | 1 | ₹950 | ₹950 | Amazon |
| 8 | Display (10.1") | HDMI IPS Touch 1280×800 | 1 | ₹5,500 | ₹5,500 | Amazon |
| 9 | Touchscreen (2.8") | ILI9341 SPI TFT | 1 | ₹650 | ₹650 | Robu.in |
| 10 | RGB LED Strip | WS2812B 1m | 1 | ₹350 | ₹350 | Amazon |
| 11 | Buzzer | 5V Active Piezo | 2 | ₹50 | ₹100 | Local |
| 12 | SD Card 64GB | SanDisk Class 10 | 1 | ₹650 | ₹650 | Amazon |
| 13 | Power Supply 5V/3A | Official RPi PSU | 1 | ₹800 | ₹800 | Robu.in |
| 14 | Enclosure | ABS Junction Box IP65 | 1 | ₹1,200 | ₹1,200 | Local |
| 15 | Misc (wires, PCB, etc) | — | — | — | ₹2,000 | Local |
| TOTAL | Team 1 Hardware Cost |  |  |  | ₹26,200 |  |

## 7.2 Team 2 – Robot BOM (Selected Items)

| # | Component | Model | Qty | Unit Price | Total | Source |
| --- | --- | --- | --- | --- | --- | --- |
| 1 | Main Computer (SBC) | Jetson Orin Nano 8GB | 1 | ₹35,000 | ₹35,000 | Arrow/E-bay |
| 2 | Safety Microcontroller | STM32F4 Discovery Board | 1 | ₹2,500 | ₹2,500 | Amazon |
| 3 | LiDAR Sensor | RPLIDAR A1M8 | 1 | ₹12,000 | ₹12,000 | Robu.in |
| 4 | Depth Camera | Intel RealSense D435i | 1 | ₹28,000 | ₹28,000 | Amazon |
| 5 | Thermal Camera | MLX90640 32×24 IR | 2 | ₹3,500 | ₹7,000 | Amazon |
| 6 | HQ Camera (Pi) | Raspberry Pi HQ Cam + 16mm | 1 | ₹8,000 | ₹8,000 | Amazon |
| 7 | UV-C Lamps (Amalgam) | 55W Germicidal UV-C | 2 | ₹3,500 | ₹7,000 | B2B Supplier |
| 8 | UV-C Lamp (Floor) | 36W UV-C Floor Lamp | 1 | ₹2,500 | ₹2,500 | B2B Supplier |
| 9 | UV-A LED Panel 365nm | 20W UV-A Panel | 1 | ₹2,800 | ₹2,800 | Amazon |
| 10 | Electronic Ballast | HF Ballast 55W UV | 2 | ₹1,800 | ₹3,600 | B2B Supplier |
| 11 | Safety Relay | Omron G9SE-201 Safety Relay | 1 | ₹4,500 | ₹4,500 | Amazon |
| 12 | DC Motor + Encoder | 12V 30RPM Planetary | 2 | ₹2,200 | ₹4,400 | Robu.in |
| 13 | Motor Driver | ODrive v3.6 or L298N | 1 | ₹6,500 | ₹6,500 | Robu.in |
| 14 | LiFePO4 Battery 24V | 24V 20Ah LiFePO4 Pack | 1 | ₹18,000 | ₹18,000 | B2B Battery |
| 15 | BMS | 24V 20A Battery Mgmt System | 1 | ₹1,500 | ₹1,500 | Amazon |
| 16 | DC-DC Converter 5V | 24V→5V 10A Buck | 2 | ₹450 | ₹900 | Robu.in |
| 17 | DC-DC Converter 12V | 24V→12V 5A Buck | 1 | ₹350 | ₹350 | Robu.in |
| 18 | Reed Switch Sensors | MC-38 Door Sensor | 8 | ₹80 | ₹640 | Robu.in |
| 19 | IMU | MPU6050 Module | 1 | ₹250 | ₹250 | Robu.in |
| 20 | Ultrasonic Sensors | HC-SR04 | 3 | ₹100 | ₹300 | Local |
| 21 | Emergency Stop Button | Red Mushroom 22mm NC | 2 | ₹350 | ₹700 | Amazon |
| 22 | V-Slot Aluminum Frame | 20×20 V-Slot – 3m lengths | 4 | ₹600 | ₹2,400 | Motedis/Local |
| 23 | Wheels + Casters | 150mm PU Wheel set | 1 | ₹2,200 | ₹2,200 | Robu.in |
| 24 | Misc Electronics | PCB, connectors, cable, etc | — | — | ₹5,000 | Local |
| TOTAL | Team 2 Hardware Cost |  |  |  | ₹1,57,340 |  |

# 8. PCB & ENCLOSURE DESIGN

## 8.1 Custom PCB Designs Required

### PCB-01: Safety Interface Board (STM32)

- Hosts STM32F4 footprint or uses dev board mounted via standoffs

- 8× optocoupler inputs for reed switch supervised loops (PC817)

- 4× relay outputs (UV lamp, buzzer, alarm light, charging relay)

- Hardware watchdog timer circuit (555 timer or dedicated IC)

- CAN bus transceiver (SN65HVD230) for inter-board communication

- 24V→5V onboard regulator (LM2596 or LM7805 heatsink)

- Design tool: KiCad 7.0 (open source). Manufacture: JLCPCB (10pcs ~$15 USD)

### PCB-02: Sensor Hub Board (Team 1)

- Arduino Mega shield with labeled I2C breakouts (×6) + UART headers (×3)

- 3.3V/5V level shifter for mixed-voltage sensors

- Screw terminals for secure wire connections in hospital environment

- Protection diodes on all I/O pins

## 8.2 Enclosure Design

### Robot Enclosure

- Frame: 20×20 V-slot aluminum extrusion – rigid, lightweight, customizable

- Panels: 3mm ABS plastic sheet (laser cut or CNC routed)

- Design software: FreeCAD or Fusion 360

- UV-C lamp housing: UV-resistant polycarbonate shield (blocks UV from wiring)

- IP rating target: IP42 (protected from solid objects >1mm + vertical dripping water)

- Cable management: UV-resistant cable sleeves + strain relief grommets

### Environment Monitoring Station Enclosure

- Wall-mounted ABS enclosure IP65 (240×190×90mm)

- Laser-engraved labels for sensor ports

- 10.1" display bezel mounted flush with front panel

- Ventilation slots with dustproof mesh on sides for air sensors

# 9. MEDICAL STANDARDS & COMPLIANCE

## 9.1 Applicable Standards

| Standard | Organization | Relevance to UV-GUARD |
| --- | --- | --- |
| ASHRAE 170-2021 | ASHRAE | OT ventilation: 20+ ACH, temp 18-24°C, humidity 45-60% |
| WHO Guidelines 2016 | WHO | Air quality in healthcare: PM2.5 <10µg/m³, NO2 <40µg/m³ |
| NABH Standards | NABH India | Hospital infection control, OT environment standards |
| IEC 60601-1 | IEC | Medical electrical equipment safety (UV lamp system) |
| ISO 14971 | ISO | Risk management for medical devices |
| EN 62471 | IEC/EN | Photobiological safety of lamps (UV hazard classification) |
| CDC Guidelines | US CDC | UV-C disinfection: 25 mW/cm² minimum, dose targets per pathogen |
| IPC-2581 | IPC | PCB design data format standard |

## 9.2 UV-C Dose Requirements by Pathogen

| Pathogen | Required Dose | Log Reduction | Time @ 25mW/cm² |
| --- | --- | --- | --- |
| C. difficile (spores) | 22,000 mJ/cm² | 3-log (99.9%) | 14.7 minutes |
| MRSA | 6,000 mJ/cm² | 3-log | 4 minutes |
| E. coli | 4,000 mJ/cm² | 3-log | 2.7 minutes |
| Acinetobacter | 7,400 mJ/cm² | 3-log | 4.9 minutes |
| SARS-CoV-2 | 3,700 mJ/cm² | 3-log | 2.5 minutes |
| Aspergillus (fungi) | 176,000 mJ/cm² | 2-log (99%) | 117 minutes* |

* Aspergillus requires chemical + UV combined approach. UV-GUARD targets C. difficile as the most resistant target (22,000 mJ/cm²), which ensures all other pathogens are eliminated.

# 10. TESTING & VALIDATION PROTOCOL

## 10.1 Unit Testing Checklist

| ID | Test Description | Expected Result | Pass/Fail | Responsible |
| --- | --- | --- | --- | --- |
| T01 | Thermal camera detects mannequin at 37°C body temp | HUMAN DETECTED within 1.5sec | [ ] | Team 2 |
| T02 | Door sensor open → UV OFF within 50ms | Relay drops <50ms measured by oscilloscope | [ ] | Team 2 |
| T03 | PMS5003 accuracy vs reference PM meter | ±15% of reference reading | [ ] | Team 1 |
| T04 | BME280 temperature accuracy | ±0.5°C vs calibrated thermometer | [ ] | Team 1 |
| T05 | CO2 sensor zero + span calibration | Read 400ppm ambient ±50ppm | [ ] | Team 1 |
| T06 | Robot navigation in 5×5m test area | Coverage >95% with no collisions | [ ] | Team 2 |
| T07 | UV dose meter at 1m from lamp | ≥25 mW/cm² at full power | [ ] | Team 2 |
| T08 | MQTT publish/subscribe latency | <100ms end-to-end | [ ] | Integration |
| T09 | Dashboard classification display | Correct color + label for all levels | [ ] | Team 1 |
| T10 | YOLOv8 organism detection (blood stain) | >80% detection rate on test panel | [ ] | Team 2 |
| T11 | Battery runtime test | >4 hours continuous operation | [ ] | Team 2 |
| T12 | Emergency stop response | All motion + UV stops <200ms | [ ] | Integration |
| T13 | Full disinfection cycle (20m² test room) | Completes without breach, report generated | [ ] | Integration |
| T14 | Post-disinfection environment classification | Correct EXCELLENT/GOOD/WARNING display | [ ] | Team 1 |

## 10.2 Integration Test Protocol

- Step 1: Run all unit tests independently on each sub-system

- Step 2: Connect Team 1 and Team 2 systems via MQTT on local WiFi. Verify all topics

- Step 3: Full-system dry run in lab without UV lamps (safety validation only)

- Step 4: Controlled UV test in shielded chamber – measure dose with calibrated radiometer

- Step 5: Full room pilot test in non-clinical environment (mock OT setup)

- Step 6: Stakeholder demonstration + feedback collection

- Step 7: Clinical pilot at partner hospital (supervised, non-patient area)

# 11. LEVEL-1 PROJECT TRACKING REPORT

Project: UV-GUARD Autonomous Disinfection System

Target Market Launch: Week 16 (4 Months from project start)

Report Level: Level 1 – Executive Tracking (Weekly Updates)

## 11.1 Phase Plan (4-Month Roadmap)

| Phase | Milestone | W1 | W2 | W3 | W4 | W5 | W6 | W7 | W8 | W9 | W10 | W11 | W12 | W13-16 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| P1 | Requirements & Design | ▓ | ▓ |  |  |  |  |  |  |  |  |  |  |  |
| P1 | Component Ordering |  | ▓ | ▓ |  |  |  |  |  |  |  |  |  |  |
| P2 | Team 1: Sensor Setup |  |  | ▓ | ▓ | ▓ |  |  |  |  |  |  |  |  |
| P2 | Team 1: HMI Dashboard |  |  |  | ▓ | ▓ | ▓ |  |  |  |  |  |  |  |
| P2 | Team 2: Chassis Build |  |  | ▓ | ▓ | ▓ | ▓ |  |  |  |  |  |  |  |
| P2 | Team 2: Nav System |  |  |  | ▓ | ▓ | ▓ | ▓ |  |  |  |  |  |  |
| P2 | Team 2: Safety System |  |  | ▓ | ▓ | ▓ |  |  |  |  |  |  |  |  |
| P2 | Team 2: UV System |  |  |  |  | ▓ | ▓ | ▓ |  |  |  |  |  |  |
| P2 | Team 2: CV Pipeline |  |  |  |  | ▓ | ▓ | ▓ | ▓ |  |  |  |  |  |
| P3 | System Integration |  |  |  |  |  |  | ▓ | ▓ | ▓ |  |  |  |  |
| P3 | Unit Testing |  |  |  |  |  | ▓ | ▓ | ▓ |  |  |  |  |  |
| P3 | Integration Testing |  |  |  |  |  |  |  | ▓ | ▓ | ▓ |  |  |  |
| P4 | Pilot Test (Mock OT) |  |  |  |  |  |  |  |  | ▓ | ▓ | ▓ |  |  |
| P4 | Fix + Optimize |  |  |  |  |  |  |  |  |  | ▓ | ▓ | ▓ |  |
| P5 | Clinical Pilot |  |  |  |  |  |  |  |  |  |  | ▓ | ▓ | ▓ |
| P5 | Market Launch Prep |  |  |  |  |  |  |  |  |  |  |  | ▓ | ▓ |

## 11.2 Milestone Tracker

| ID | Milestone | Target Date | Status | Owner / Notes |
| --- | --- | --- | --- | --- |
| M01 | All components ordered and received | Week 2 | 🔵 Planned | All teams / Long-lead items: Jetson, RealSense |
| M02 | Team 1 sensor readings verified (all 6 sensors) | Week 5 | 🔵 Planned | Team 1 Lead |
| M03 | Team 1 HMI dashboard v1.0 running | Week 6 | 🔵 Planned | Team 1 Dev |
| M04 | Team 2 chassis assembled + motors running | Week 5 | 🔵 Planned | Team 2 Mech |
| M05 | SLAM mapping in 5×5m test area working | Week 7 | 🔵 Planned | Team 2 Nav |
| M06 | Safety FSM + thermal + door interlock verified | Week 6 | 🔵 Planned | Team 2 Safety – CRITICAL |
| M07 | UV-C system integrated + dose measured | Week 7 | 🔵 Planned | Team 2 UV |
| M08 | YOLOv8 model trained (mAP50 >0.75) | Week 8 | 🔵 Planned | Team 2 CV |
| M09 | MQTT integration all topics live | Week 8 | 🔵 Planned | Integration |
| M10 | Full unit test suite passed (14 tests) | Week 9 | 🔵 Planned | All teams |
| M11 | Full integration test in mock lab | Week 10 | 🔵 Planned | Integration Lead |
| M12 | Pilot test in 20m² mock OT – full cycle | Week 11 | 🔵 Planned | All teams |
| M13 | Clinical pilot (partner hospital) | Week 14 | 🔵 Planned | Project Lead + Hospital Liaison |
| M14 | Market launch documentation + demo video | Week 16 | 🔵 Planned | All teams |

## 11.3 Risk Register

| ID | Risk | Probability | Impact | Mitigation Strategy |
| --- | --- | --- | --- | --- |
| R01 | Jetson Orin Nano delivery delay (import) | HIGH | HIGH | Order immediately. Have Raspberry Pi 5 as backup. Port Nav+CV to RPi5 if needed. |
| R02 | YOLOv8 organism detection accuracy <75% | MEDIUM | HIGH | Collect more training data. Use ensemble model. Add UV-A intensity calibration. |
| R03 | Thermal camera false positives (hot equipment) | MEDIUM | HIGH | Fine-tune detection algorithm. Create exclusion zones. Use temporal filtering. |
| R04 | UV-C lamp integration compliance / safety | LOW | VERY HIGH | Follow IEC 60601-1. Hardware interlock. Safety review before any UV test. |
| R05 | MQTT communication latency in safety path | LOW | HIGH | Safety-critical signals use hardware (GPIO/relay) NOT MQTT. MQTT is telemetry only. |
| R06 | Hospital regulatory approval for clinical pilot | MEDIUM | MEDIUM | Engage hospital biomedical team early. Prepare IEC 60601 documentation. |
| R07 | Budget overrun on robot hardware | MEDIUM | MEDIUM | Contingency 20% buffer. Substitute Jetson with RPi5+Coral TPU if needed. |
| R08 | Team bandwidth / skill gaps | MEDIUM | MEDIUM | Weekly sync meetings. Shared GitHub repo. Pair programming for complex modules. |

## 11.4 Weekly Status Template

Each team lead fills this template every Friday and shares in group chat:

=== UV-GUARD WEEKLY STATUS – TEAM [X] – WEEK [N] ===
📅 Date: ___________
✅ Completed this week:
   1. ___
   2. ___
🔄 In Progress:
   1. ___
⚠️  Blockers / Issues:
   1. ___
📋 Plan for next week:
   1. ___
📊 Milestone Status: [ON TRACK / AT RISK / DELAYED]
💸 Budget used this week: ₹___

# 12. BUDGET ESTIMATION

## 12.1 Total Project Cost Summary

| Budget Category | Estimated Cost | Notes |
| --- | --- | --- |
| Team 1 – Environment Monitoring Hardware | ₹26,200 | All sensors + RPi4 + display |
| Team 2 – Robot Hardware (Mechanical) | ₹28,000 | Frame + motors + wheels + battery |
| Team 2 – Robot Electronics & Sensors | ₹1,29,340 | Jetson + RealSense + LiDAR + UV lamps |
| PCB Manufacturing (JLCPCB – 2 designs) | ₹8,000 | 10pcs each + assembly |
| Enclosures + Fabrication | ₹12,000 | ABS panels, CNC, 3D printing |
| Optical Components (UV filter, lenses) | ₹5,000 | Long-pass filter + camera lens |
| Testing Equipment | ₹8,000 | UV radiometer rental + calibration |
| Software Licenses & Cloud | ₹3,000 | 4 months cloud hosting |
| Miscellaneous + Contingency (20%) | ₹40,000 | Cables, tools, replacements |
| TOTAL PROJECT BUDGET | ₹2,59,540 (~$3,100 USD) | Prototype / MVP |

## 12.2 Market Pricing Estimate (Product Launch)

- Target selling price per unit (hospital): ₹8,00,000 – ₹12,00,000

- Annual maintenance contract: ₹80,000 – ₹1,20,000/year

- Disinfection-as-a-Service model: ₹15,000–25,000 per disinfection session

- Target customers: Private hospitals, government hospital OTs, diagnostic centers, clean rooms

## 12.3 Next Steps for Market Launch

- File for startup registration (DPIIT Startup India recognition)

- Apply for CDSCO (Central Drugs Standard Control Organisation) approval for medical device

- File provisional patent for autonomous UV safety interlock system

- Partner with 2–3 hospitals for paid clinical pilot program

- Develop marketing materials: demo video, brochure, website

- Approach hospital equipment distributors for distribution agreement

UV-GUARD  |  Protecting Lives Through Intelligent Disinfection  |  Version 1.0  |  April 2026